MP-100

MICROPROCESSORIZED INSTRUMENT FOR DETERMINATION OF THE RELATIVE MAGNETIC PERMEABILITY OF NON-MAGNETIC MATERIALS BY NON-DESTRUCTIVE METHOD

|  | Mobile meter designed to measuring the magnetic permeability lower as 3 and allows to control the elements with small overall dimensions and the elements with complex shapes as well as it allows the evaluations of heterogeneity of the elements with big overall dimensions. The meter is internal battery-powered what allows about 12h continuous work. |
| --- | --- |

The device is fully automated and it does not require any regulations, including charging a battery. The meter contain the measurements memory and all obtained date could be display on LCD.

On request the instrument enables the evaluation of welds of non-magnetic steel in an aspect of magnetic.

The MP-100 meter provides the opportunity of a binary and graphic presentation of the obtained data, which can be presented by:

- LCD display,

# EQUIPMENT

The MP-100 meter contains:

- Sensor,

- Reference standard of material,

- Certificate of reference standard of material,

- RS Cable,

- Cover,

- Supply unit,

- DTR – 2 pieces.

# TECHNICAL DATA

Measurement range					1 ( 2.99

Resolution						0.01

Digital read-out	alphanumeric display LCD

Measurement accuracy

(measurement on the standard sample)		( 5 %

Signalling of range exceeding				graphic and acoustic

Number of saved measurement results		ca. 1200 results for 30 objects

Determination and presentation measured values 
of permeability for each object as: 			minimum,

maximum,

average

Auto-calibration of the device			Yes

Bar-graph mode					Yes

Data transmission to the computer			series interface

The remembered results presentation: 	LCD display or PC monitor (after meter connection with PC)

The printing of measurements report by:	serial printer directly from MP-100 meter or PC printer (after meter connection with PC)

Deleting unnecessary results

Choosing of language version: 			polish/english

Control						Keyboard

Supply							Accumulator

Signalling of necessity of charging a battery		acoustic and graphic

Operating conditions:

- ambient temperature				0(C ( + 50(C

- air relative humidity				20 % ( 80 %

# DIMENSION

all dimensions in mm